## Slide 1
James Tolchard
NMR Laboratory Monitoring
24th August 2023

## Slide 2
Why ?
- Network Connectivity Issues
- Quantitative Evaluation of Helium Recycling

## Slide 3
Inspiration
Resource Dashboard
Helps troubleshooting
  - Problems / correct departments
Long-term trends
Metric triggers
  Preemptive-intervention
  Alerts:
    Automated comms. (Email/Slack)
    Scriptable reactions
Containerization

## Slide 4
How ? - Containerization
Scalable
Cross-platform
Low Overheads
Fast
Free
Easy to develop
Key requirements
Infrastructure
Host OS
Hypervisor
Guest OS
Guest OS
Bins / Libs
Bins / Libs
App 1
App 1
Virtual Machine
Infrastructure
Host OS
Container Engine
Bins / Libs
Bins / Libs
App 1
App 1
Container
Container Engine

## Slide 5
How ? - Containerization
Simple / User friendly CLI
Easy to Develop
Straightforward local installation
Community / Documentation / Support
Open source (core) & free to use
Vast Eco-system (100,000+ images)
Container Engines / Platforms
+ more…

## Slide 6
SCANS
Simple Containerized Analysis for NMR Systems
Containers
Time-Series database
Scrapes data over HTTP
Fixed retention time per DB
CRMN Server
  Cryogen metrics (low res., 2-year retention)
  Computer metrics (high res. 2-week retention

## Slide 7
SCANS
Simple Containerized Analysis for NMR Systems
Spectrometer containers
CRMN Server
Collects computer metrics
Node Exporter
Checks HLMU readings
Non-MICS cryogens
Collects MICS metrics
Spectrometer

## Slide 8
SCANS
Simple Containerized Analysis for NMR Systems
Web-scraping containers
CRMN Server
Accesses Air Liquide APIs for Tank & SPI information
Monitors external conditions via API
External weather

## Slide 9
SCANS
Simple Containerized Analysis for NMR Systems
Hardware-scraping containers
CRMN Server
Python script creates logs via ModBus comm.
Compressor info
Scrapes UPS details via local web-interface
UPS info
Temp, humidity & pressure via custom Arduino API
NMR Hall Env. sensor

## Slide 10
SCANS
Simple Containerized Analysis for NMR Systems
Hardware-scraping containers
CRMN Server
RAID Storage metrics
Dell Server Metrics
Analyses network response times
Blackbox exporter

## Slide 11
SCANS
Simple Containerized Analysis for NMR Systems
Visualization & analysis container
CRMN Server
Unique user logins and roles (i.e., admin, viewers)
Create alerts based on metrics and send custom notifications
Query Prometheus DBs and create interactive charts and dashboards

## Slide 12
SCANS
SCANS on GitHub (currently private)

## Slide 13: SCANS Dashboard: Summary

## Slide 14: SCANS Dashboard: Cryogen metrics

## Slide 15: SCANS Dashboard: Spectrometer metrics

## Slide 16: SCANS Dashboard: Workstation metrics

## Slide 17: SCANS: Current Automated Alerts
UPS Discharging
Server room air conditioning fail (>30°C)
RAID disk fails
Compressor room air conditioning fail (>30°C)
Helium full – needs changing (>195 bar)
FLOXAL N2 low pressure (<1 bar)
Air Liquide LN2 tanks, high usage (20%/hr)
Air Liquide LN2 tanks, low level (<33%)
Alerts currently sent to custom #alerts Slack & Teams channels

## Slide 18
Observations so far…
Normally, NMR hall temp is regulated between ~20-21°C)

## Slide 19
Observations so far…
But, in a heat wave, the Air Conditioning regulation fails… and NMR Hall can reach ~30.4°C!
Spec. | ∆ Shim Temp (°C)
400 | +8
500 | +7
600 | +7
800nb | -3
800wb | +8
GHz | +11
F80 | ? (but lock lost)

## Slide 20
Observations so far…
Humans make mistakes…
- new helium rack connected (but not opened)
26 Hrs

## Slide 21
Observations so far…
Our Network connectivity issues seem to be resolved

## Slide 22
Observations so far…
The Topspin data directories are mounted incorrectly on the 500 MHz and 800 MHz WB
- luckily informed user before summer experiments!